Objective: Quantify HiBit-tagged protein

Materials:

- Cell free lysate containing HiBit-tagged protein

- LgBit stock solution (53.2 μM) (this exact concentration is not important; the working concentration is a 6x dilution (8-9 μM))

- Coelenterazine H (CTZ-H) stock solution (10 mM in DMSO)

- 10 mM Tris pH=7.5

- 384 well plate with square white bottom (96 well plate works too, if you have a standard curve prepared)

Dilutions:

| Dynamic Range: | Dilution Factor: | Source: | Source Volume (μL): | Diluent Volume (μL): |
| --- | --- | --- | --- | --- |
| 25-100 μg/mL | 50x | Stock | 2 | 98 |
| 100-400 μg/mL | 200x | 50x | 20 | 60 |
| 400-1600 μg/mL | 800x | 200x | 20 | 60 |

*10 mM Tris pH=7.5 is the diluent

The optimal range for quantification is between 0.5-2 μg/mL, so you will want at least one of your protein dilutions to hit within this range. The above dilution factors are likely to achieve this range for most CFE expression ranges, but you may have to tweak this based on your own levels of protein expression.

Protocol:

| X | Step | Value |
| --- | --- | --- |
|  | Thaw LgBit stock on ice while preparing dilutions | See table |
|  | Prepare 6x LgBit dilution (LgBit stock ¦ Tris) (1 stock tube per 50 reactions) | 20 μL ¦ 100 μL |
|  | Prepare CTZ-H working reagent (40 μM) (CTZ-H stock ¦ Tris) (50 reactions) | 1 μL ¦ 249 μL |
|  | Transfer samples to 384 well plate (square white bottom) on ice | 2 μL |
|  | Transfer LigBit to 384 well plate on ice | 2 μL |
|  | Spin plate down | - |
|  | Transfer CTZ-H to 384 well plate with mixing | 4 μL |
|  | Spin plate down | - |
|  | Run plate reader protocol | 30 minutes |
|  | Use HiBitQuant to quantify | - |

*Note: you want to do many of these steps on ice to slow down the reaction’s initiation. You want to minimize the time between adding CTZ-H and starting the plate reader

*You may also need to tweak the plate reader protocol for its kinetic interval if the interval is too short for the number of reads you need to do

Plate Reader Protocol:

- Orbital: 0:05 (MM:SS)

- Frequency: 365 cpm (2 mm)

- Runtime 0:30:00 (HH:MM:SS), Interval 0:01:53, 21 Reads

- Luminescence Endpoint

- Integration Time: 0:01.00 (MM:SS.ss)

- Filter Set 1

Emission: Full light

Optics: Top,  Gain: 235

- Read Speed: Normal,  Delay: 100 msec

- Extended Dynamic Range

- Read Height: 1 mm